REAGEN™

Nitrate/Nitrite Assay Kit

RNF91004

Product Description

REAGEN™ Nitrate/Nitrite Assay Kit is a cuvette-based enzymatic assay for the determination of Nitrate/Nitrite across a variety of food samples. The kit uses a spectrophotometric, kinetic assay to detect the enzymatic degradation of Nitrate/Nitrite directly from food samples.

The unique features of the kit are:

- High sensitivity and low detection limit (0.02 ppm)

- High reproducibility

- A rapid (45 minute), robust enzyme-based assay

Procedure Overview

REAGEN™ Nitrate/Nitrite Assay Kit uses a coupled enzymatic reaction scheme to detect nitrate and/or nitrite in food samples: nitrate in the sample is converted by nitrate reductase enzyme to make nitrite. The nitrite then reacts with sulfanilamide and N-(1-naphthyl)-ethylene-diamine to form a red diazo dye. This dye, measured at 540 nm, is proportional to the level of nitrate and/or nitrite in the sample.

The absorbance of each sample at 540 nm is measured using a spectrophotometer. The concentration of nitrate /nitrite in each sample is then directly determined from the diazo dye absorbance at 540 nm within 45 minutes.

Kit Contents, Storage and Shelf Life

REAGEN™ Nitrate/Nitrite Assay Kit has the capacity for 48 determinations or 24 determinations in duplicate. Store the kit at 4°C. The shelf life of the kit is 12 months, after receipt, when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Buffer Solution | 4 mL | 4°C |
| Reagent Mix | 1 vial | 4°C |
| Nitrate Reductase | 1 vial | 4°C |
| Color Reagent 1 | 3 mL | 4°C |
| Color Reagent 2 | 3 mL | 4°C |
| Microtiter Plate | One | 4°C |

Sensitivity (Detection Limit)

| Sample Volume | Detection Limit (mg/L) |
| --- | --- |
| 0.1 mL | 0.02 |

The linear range of the assay is 0.02 –40 mg/L (ppm).

Required Materials Not Provided With the Kit

- Spectrophotometer (578 nm)

- Low volume (0.3 mL) UV cuvettes, such as Brand Scientific catalog #7592 00

- Deionized or distilled water

- 100 mL volumetric flask

- 50 mL plastic conical tubes

- Carrez buffer I and Carrez buffer II (recipes included in protocol)

Warnings and Precautions

- Do not use the kit past the expiration date.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays  in  direct sunlight,  as this may cause excessive  heat and evaporation. Cold  bench tops should be avoided by placing several layers of paper towel or some other insulate material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents  into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

SAMPLE PREPARATION

### General information on sample preparation

- Use clear, colorless or faintly colored solutions directly of dilute to a concentration of <5 mg/L nitrite or nitrate before using in test

- Filter turbid solutions or clarify using Carrez solutions and dilute such that the nitrite or nitrate concentration is below 5 mg/L. Diluted solution can be directly used in the test.

- After adding Carrez solutions I and II it is important to adjust pH to 8.0 ± 0.2.

### Water, Waste Water, Drinking Water and Mineral Water

If necessary, dilute clear, colorless water samples to a nitrite or nitrate concentration of <5 mg/L and use in the test. Filter turbid water samples or samples containing carbon dioxide (to remove the gas) before use in the test.

### Fruit Juices

Accurately weigh 2.5 g of homogenized sample into a 50 mL conical tube and add 10 mL dH2O and mix. Add 2.5 mL diluted Carrez solution (3.6 g potassium hexacyanoferrate[II] trihydrate/100 mL), mix well. Add 2.5 mL Carrez solution II (7.2 g zinc sulfate heptahydrate/100 mL) and mix well again. Adjust pH to 8.0 with 1M sodium hydroxide solution. Bring solution up to 50 mL with dH20 and mix well. Filter if necessary or centrifuge to clarify for testing. Use the clear liquid for the test.

### Dairy Products

Accurately weigh out 2 g sample and transfer to a 100 mL volumetric flask. Add 50 mL of boiling dH20 and mix well. Boil on a water bath for 15 min and then allow sample to cool to ~ 25°C.  Add 3 mL each concentrated Carrez I (15 g potassium hexacyanoferrate[II] trihydrate/100 mL) and Carrez II (30 g zinc sulfate heptahydrate/100 mL) solutions, mixing well. Adjust pH to 8.0 with 1M sodium hydroxide solution. Bring solution up to 100 mL with dH20 and mix well. Transfer an aliquot 10 - 20 mL to a tube and centrifuge at 3,000 x g for 15 minutes. Filter the supernatant with a membrane filter and use the clear liquid for the test.

### Meat Products

Accurately weigh out 5 g of well homogenized sample into a 100 mL beaker and add 50 mL of boiling dH20. Boil on a water bath for 15 min and then allow sample to cool to ~ 25°C. Add 3 mL each concentrated Carrez I (15 g potassium hexacyanoferrate[II] trihydrate/100 mL) and Carrez II (30 g zinc sulfate heptahydrate/100 mL) solutions, mixing well. Adjust pH to 8.0 with 1M sodium hydroxide solution. Bring solution up to 100 mL with dH20 and mix well. Transfer an aliquot 10 - 20 mL to a tube and centrifuge at 3,000 x g for 15 minutes. Filter the supernatant with a membrane filter and use the clear liquid for the test.

### Potato-based Foods, Fruits and Vegetables

Accurately weigh out 3 g of well homogenized sample into a 100 mL beaker and add 60 mL of hot dH20. Mix well or shake, and incubate for 15 min. on a water bath set at 70°C. Treat samples that are intensely colored or that have high starch content with Carrez solutions (same concentration as the fruit juice protocol listed above). Allow contents of the beaker to cool to ~ 25°C, and bring volume of solution up to 100 mL with dH20.  Filter if necessary and use the clear liquid for the test.

### Cheese and Processed Cheese Products

Accurately weigh out 3 g of homogenized sample into a 100 mL volumetric flask, add 20 mL dH20 and heat in microwave oven until the cheese is melted (about 10 - 20 sec). Stir with a spatula and microwave again until the mixture boils gently. Add about 50 mL of boiling dH20 and boil for 15 min on a water bath. Cool solution to ~ 25°C and add 5 mL each of concentrated Carrez I (15 g potassium hexacyanoferrate[II] trihydrate/100 mL) and Carrez II (30 g zinc sulfate heptahydrate/100 mL) solutions, mixing well. Adjust pH to 8.0 with 1M sodium hydroxide solution. Bring solution up to 100 mL with dH20 and mix well. Transfer an aliquot 10 - 20 mL to a tube and centrifuge at 3,000 x g for 15 minutes. Filter the supernatant with a membrane filter and use the clear liquid for the test.

### Baby Food Products

Accurately weigh out 1.5 g of well homogenized sample into a 100 mL beaker and add 50 mL of boiling dH20. Boil on a water bath for 15 min and then allow sample to cool to ~ 25°C. Add 3 mL each concentrated Carrez I (15 g potassium hexacyanoferrate[II] trihydrate/100 mL) and Carrez II (30 g zinc sulfate heptahydrate/100 mL) solutions, mixing well. Adjust pH to 8.0 with 1M sodium hydroxide solution. Bring solution up to 100 mL with dH20 and mix well. Transfer an aliquot (10 - 20 mL) to a tube and centrifuge at 3,000 x g for 15 minutes. Filter the supernatant with a membrane filter and use the clear liquid for the test.

NITRATE/NITRITE DETECTION PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at20 – 25°C / 68 – 77°F); Make sure you read “Warnings and Precautions” section on page 3.

Prepare Nitrate Reductase: Reconstitute the lyophilized Nitrate Reductase by adding 275 µL dH20 and gently swirling to mix.

Prepare Reagent Mix solution: Reconstitute the Reagent Mix by adding 3 mL Buffer Solution to the vial containing a Reagent Mix. Gently swirl or invert the vial to completely dissolve the tablet.

Prepare 500 mg/L sodium nitrite standard: Weigh 75 mg sodium nitrite into a 100 mL volumetric flask, dissolve in dH2O and make up to the mark with the same solvent (stock solution = 500 mg nitrite/L). Prepare a nitrite calibration curve by serially diluting the stock solution with dH2O to give concentrations in the range from 5.00 mg and 0.05 mg nitrite/L.

Prepare 500 mg/L potassium nitrate standard: Weigh 81.5 mg potassium nitrate into a 100 mL volumetric flask, dissolve in dH2O and make up to the mark with the same solvent (stock solution = 500 mg nitrite/L). Prepare a nitrate calibration curve by diluting the stock solution with dH2O to give concentrations in the range from 5.00 mg and 0.05 mg nitrate/L.

### ELISA Testing Protocol

Note: The determinations are performed in 0.3 mL microcuvettes. To ensure that the color change in the assay is specifically produced by Nitrate/Nitrite, each determination is performed, in parallel, using samples treated with enzyme (ascorbate oxidase) to degrade Nitrate/Nitrite (Sample blank). The color change in the Nitrate/Nitrite depleted reaction (Sample blank) is then subtracted from the color change of corresponding untreated reactions (Sample reaction) to correct for individual sample background signals.

| Pipet into low volume cuvettes | Nitrite blank | Nitrite sample | Nitrite +nitrate blank | Nitrite + nitrate sample |
| --- | --- | --- | --- | --- |
| Sample or standard dH20 Reaction Mix solution Nitrate Reductase | ---- 150 μL ---- ---- | 10 - 100 μL Up to 150 μL ---- ---- | ---- 100 μL 45 μL 5 μL | 10 – 100 μL Up to 100 μL 45 μL 5 μL |
| Mix and incubate 10 min at 37°C or 15min at room. Measure the absorbance at 540 nm (A1). |  |  |  |  |
| Color reagent 1 Color reagent 2 | 50 μL 50 μL |  |  | 50 μL 50 μL |
| Mix and incubate 15 min at room temperature. Measure the absorbance at 540nm (A2). |  |  |  |  |

### Determination of Nitrate/Nitrite in Food Samples

To calculate the nitrate and nitrite concentrations, use the results calculated from the calibration curves constructed using the standard solutions. Plot the change in absorbance obtained for the sodium nitrite and potassium nitrate standard solutions on the y-axis against the corresponding nitrite or nitrate concentrations in mg/l on the x-axis.

ΔAnitrite = (A2 – A1)nitrite sample – (A2 – A1)nitrite blank

ΔAnitrite+ nitrate = (A2 – A1)nitrite + nitrate sample – (A2 – A1)nitrite + nitrate blank

ΔAnitrate = ΔAnitrite+ nitrate - ΔAnitrite

Determine the concentrations of nitrite and nitrate in the sample from the calibration curves using the change in absorbance measured.

Note: The calibration curves do not have to be plotted every time a determination is made. It is sufficient to check the calibration curves from time to time and to include a sodium nitrite or potassium nitrate standard solution as a control.

If the sample has been diluted during the sample preparation protocol, the result must be multiplied by the dilution factor (F).

When analyzing solid or semisolid samples that have to be weighed out; calculate the result with respect to the mass of sample.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250        FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.:www.reagen.us